## Slide 1: Closed Conduit Measurement Techniques
Pipeline systems
  Transmission lines
  Pipe networks
  Measurements
  Manifolds and diffusers
  Pumps
  Transients
You are here

## Slide 2: Measurement Techniques
Direct Volume or Weight measurements
Velocity-Area Integration
Pressure differential
  Pitot Tube
  Venturi Meter
  Orifice
  Elbow Meter
Electromagnetic Flow Meter
Turbine Flow Meter
Vortex Flow Meter
Displacement Meter
Ultrasonic flow meter
Acoustic Doppler
Laser Doppler
Particle Tracking

## Slide 3: Some Simple Techniques...
Direct Volume or Weight measurements
  Measure volume and time (bucket and stopwatch)
  Excellent for average flow measurements
Velocity-Area Integration
Stream flow

## Slide 4: Pitot Tube
V
V1 =
1
2
Connect two ports to differential pressure transducer. Make sure Pitot tube is completely filled with the fluid that is being measured.
Solve for velocity as function of pressure difference
z1 = z2
Static pressure tap
Stagnation pressure tap
0

## Slide 5: Venturi Meter
1797 - Venturi presented his work on the Venturi tube
1887 - first commercial Venturi tube produced by Clemens Herschel
Minimal pressure loss
1
2
Bernoulli equation applicable?_______ Why?
Yes!
Contraction

## Slide 6: Venturi Meter Discharge Equation
Cv is the coefficient of velocity. It corrects for viscous effects (energy losses) and velocity gradients (a).
Kventuri is 1 for high Re and small D2/D1 ratios
1
2

## Slide 7: Orifice
The flow coefficient, Korifice, is a function of the ratio of orifice diameter to pipe diameter and is a weak function of ________ number.
2.5 D
8 D
D
Dh
Reynolds

## Slide 8: Elbow Meter
Acceleration around the bend results in higher pressure at the outside of the bend
Any elbow can be used as the meter
Needs to be calibrated (no standard calibration curves are available)

## Slide 9: Electromagnetic Flow Meter
Conductor moving through a magnetic field generates an _______ field.
Voltage is proportional to velocity
Causes no __________ resistance to flow
High signal amplification is required
magnet
electrodes
conductive fluid
measure voltage here
electric
“measurable”

## Slide 10: Turbine and Paddle Wheel Flow Meters
Simply a turbine mounted in a pipe held in a stream
The angular velocity of the turbine is related to the velocity of the fluid
Can operate with relatively low head loss
Needs to be calibrated
Used to measure _________ ___ ____ or ___________
volumetric flow rate
velocity

## Slide 11: Vortex Flow Meter
Vortex shedding
Strouhal number, S, is constant for Re between 104 and 106
Vortex shedding frequency (n) can be detected with pressure sensors
d
L

## Slide 12: Displacement Meter
Used extensively for measuring the quantity of water used by households and businesses
Uses positive displacement of a piston or disc
Each cycle of the piston corresponds to a known volume of water
Designed to accurately measure slow leaks!

## Slide 13: Ultrasonic Flow Meters: Doppler effect
Transmitted frequency
  The transmitted frequency is altered linearly by being reflected from particles and bubbles in the fluid. The net result is a frequency shift between transmitter and receiver frequencies that is proportional to the velocity of the particles.
http://www.sensorsmag.com/articles/1097/flow1097/main.shtml
Doppler shift
Sound velocity

## Slide 14: Ultrasonic Flow Meters: Transit Time
Measure the difference in travel time between pulses transmitted in a single path along and against the flow.
Two transducers are used, one upstream of the other. Each acts as both a transmitter and receiver for the ultrasonic beam.

## Slide 15: Acoustic Doppler Velocimeter
http://www.sontek.com/
_______ measurement
Point

## Slide 16: Laser Doppler Velocimetry
http://www.tsi.com/
a single laser beam is split into two equal-intensity beams which are focused at a point in the flow field.
An interference pattern is formed at the point where the beams intersect, defining the measuring volume.
Particles moving through the measuring volume scatter light of varying intensity, some of which is collected by a photodetector.
The resulting frequency of the photodetector output is related directly to particle velocity.
_______ measurement
Point

## Slide 17: Particle Tracking Velocimetry
http://amy.me.tufts.edu/
velocity field
Illuminate a slice of fluid (seeded with particles) with a laser sheet
Take a high resolution picture with a digital camera
Repeat a few milliseconds later
Compare the two images to determine particle displacement
Measures _______ ______

## Slide 18: Questions to Ponder
Will an ADV need to be recalibrated if it is moved from freshwater to saltwater?
A graduate student proposes to use an LDV in a wave tank (through a glass bottom) that is stratified with freshwater on top of saltwater to measure turbulence from the breaking waves. What problems might arise?
How could the flow normal to the plane of the light sheet be estimated using PTV?
Would it be possible to know the direction of the flow in the 3rd dimension?

## Slide 19: More Questions to Ponder
Why would a flow meter manufacturer specify that the pipe used for installing the meter must be straight for 10 diameters upstream and 5 diameters downstream from the meter?
How could an ultrasonic device get information about velocity at more than one location without moving (profiling)?
How could you apply the results from profiling to improve the flow rate measurement in a pipe?

## Slide 20: Orifice Example
Estimate the orifice diameter that will result in a 100 kPa pressure drop in a 6.35 mm I.D. pipe with a flow rate of 80 mL/s. The orifice coefficient (Korifice) is 0.6.
What is  the ratio of orifice diameter to pipe diameter?
If the smallest pressure differential that can accurately be measured with the pressure sensor is 1 kPa, what is the smallest flow that can accurately be measured using this orifice?
What are two ways of extending the range of measurement to lower flows?

## Slide 21: Orifice Solution
Estimate the orifice diameter that will result in a 100 kPa pressure drop in a 6.35 mm I.D. pipe with a flow rate of 80 mL/s. The orifice coefficient (Korifice) is 0.6.

## Slide 22: Orifice Solution
What is  the ratio of orifice diameter to pipe diameter?
If the smallest pressure differential that can accurately be measured with the pressure sensor is 1 kPa, what is the smallest flow that can accurately be measured using this orifice?
What are two ways of extending the range of measurement to lower flows?
8 mL/s
(0.546)